## List_of_TE
Příloha č. 2: Plnění požadovaných technických parametrů
Název zakázky | Vybavení mobilní kalibrační laboratoře - 1.část
Druh zakázky | Dodávka strojního vybavení
Místo plnění zakázky | OPTOKON a.s., Červený Kříž 250, 586 01 Jihlava
Uvedené technické parametry jsou stanoveny jako minimální požadavky. Uchazeč je oprávněm nabídnout řešení s lepími parametry.
Požadované technické parametry |  | Splňuje / hodnota
Optický attenuátor SM | Insertion loss at 0 dB ≤ 0.9 dB
Polarization-dependent loss ≤ 0.08 dB
Return loss ≥55 dB
Maximum input power +23 dBm
Wavelength range 1260 to 1650 nm
Attenuation range 70 dB
Shutter isolation ≥80 dB
Attenuation flatness ± 0.04 dB
Attenuation slew rate ≥25 dB/s
Relative attenuation uncertainty ± 0.1 dB
Attenuation repeatability ± 0.01 dB
Attenuation resolution 0.001 dB
Attenuation settling time ≤55 ms
Warm-up time 30 min
Calibration period 3 years
Operating temperature 0 to 50°C
Storage temperature –30 to 60°C
Operating humidity 15 to 80% RH, 0 to 40°C noncondensing
Optický attenuator MM | Insertion loss at 0 dB ≤ 1.5 dB
Return loss ≥45 dB
Maximum input power +23 dBm
Wavelength range 750 to 1350 nm
Attenuation range 65 dB
Shutter isolation ≥75 dB
Attenuation slew rate ≥20 dB/s
Relative attenuation uncertainty ± 0.1 dB
Attenuation repeatability ± 0.01 dB
Attenuation resolution 0.001 dB
Attenuation settling time ≤ 55 ms
Warm-up time 30 min
Calibration period 3 years
Operating temperature 0 to 50°C
Storage temperature –30 to 60°C
Operating humidity 15 to 80% RH, 0 to 40°C noncondensing
Optický switch SM | Wavelength range 1250 to 1650 nm
Insertion loss (IL) 0.7 dB
Return loss (RL) 62 dB
Polarization-dependent loss (PDL) 0.04 dB
Repeatability Sequential switching ±0.005 dB
Repeatability Random switching ±0.025 dB
IL stability (maximum) ±0.025 dB
Crosstalk (maximum) –80 dB
Max input power (optical) 300 mW
Lifetime 100 million switching cycles
Switching time Electro-mechanical (break to make) 20+10*(N–1) ms
Switching time Settling time to 90% final IL 60 ms
Switching time Settling time for 99% final IL 90 ms
Operation temperature 0 to 50°C
Operation humidity 15 to 80% RH, 0 to 40°C noncondensing
Storage temperature –30 to 60°C
Optický switch MM | Wavelength range 760 to 1360 nm
Insertion loss (IL) 0.9 dB
Return loss (RL) 40 dB
Repeatability Sequential switching ±0.005 dB
Repeatability Random switching ±0.025 dB
IL stability (maximum) ±0.025 dB
Crosstalk (maximum) –60 dB
Max input power (optical) 300 mW
Lifetime 100 million switching cycles
Switching time Electro-mechanical (break to make) 20+10*(N–1) ms
Switching time Settling time to 90% final IL 60 ms
Switching time Settling time for 99% final IL 90 ms
Operation temperature 0 to 50°C
Operation humidity 15 to 80% RH, 0 to 40°C noncondensing
Storage temperature –30 to 60°C
Paltforma | Capacity 8 modules
CPU Power PC architecture
Operating System Linux
Internal Storage 200 MB user flash storage
Remote interface USB, GPIB, Ethernet 10/100/1000Base-T
USB device compatibility Mouse, keyboard, memory stick
Display External MAP-200KD (optional) or standard DVI monitor
USB host ports 2 rear and 2 front
USB device ports 1 front
LAN 1 rear
GPIB 1 rear
DVI video 1 rear
LXI triggers 25-pin Micro-D connectors
Driver type IVI-compliant
Driver compatibility LabVIEW, LabWindows, Visual C++, Visual Basic
Accessibility Multi-user sharing support
Power 100 to 250 V AC, 50/60 Hz, Auto-switching (field-replaceable as part of the power supply controller module)
Power consumption 285 VA
Local interlock Key located in front
Remote interlock Terminals located in rear
Operating temperature 0 to 50°C
Storage temperature –30 to 60°C
Humidity 15–80% RH, 0 to 40°C noncondensing
Mikroskop | Live image 800 x 600, 15 fps
Connector USB 2,0
Cord length 183 cm (6 ft)
Camera sensor 1280 x 1024 black and white, 1/3 in (1,27 cm) CMOS
Particle size detection < 0,5 um
Light source Blue LED, 100000+ hour life
Lighting technique Coaxial
Power source USB port
Certification CE
Field of View values - High Mag. (um): 185 Horizontal, 140 Vertical, 230 Diagonal
Field of View values - High Mag. (um): 300 Horizontal, 225 Vertical, 375 Diagonal